## Slide 1: “Absolute” Calibration of PMT
University of California, Los Angeles
Department of Physics and Astronomy
arisaka@physics.ucla.edu
Katsushi Arisaka

## Slide 2: Talk Outline
Introduction: Principle of PMT
NIST Standard Silicon Photodiode
Uncertainties, Concerns Specific to PMTs
Suggestions and Proposal to our Community

## Slide 3: Talk Outline
Introduction: Principle of PMT
NIST Standard Silicon Photodiode
Uncertainties, Concerns Specific to PMTs
Suggestions and Proposal to our Community

## Slide 4: Emission, Propagation and Detection of Fluorescent Photons
Emission
Propagation
Air Fluorescence Detector
Detection

## Slide 5: Emission Spectrum of Nitrogen Fluorescence
From HiRes
313nm
337nm
357nm
391nm

## Slide 6: Photon Detection Efficiency of Auger FD
GAP 2002-029
Mirror
Schmidt Corrector
UV-Filter
PMT QE
Convoluted Efficiency

## Slide 7: Principle of PMT
How PMT Works
Fundamental Parameters of PMT
  Quantum Efficiency (QE)
  Photoelectron Collection Efficiency (Col)
  Gain (G)
  Excess Noise Factor (ENF)
  Energy Resolution (/E)
How to Measure These Parameters
Some Remarks

## Slide 8: Structure of Linear-focus PMT
Mesh Anode
Last Dynode
Photo Cathode
Second Last
Dynode
First Dynode
Glass Window
Photons
QE
Col
1
2
3
n
N

## Slide 9: Quantum Efficiency (QE)
Definition:
How to measure:
  Connect all the dynodes and the anode.
  Supply more than +100V for 100% collection efficiency.
  Measure the cathode current (IC).
  Compare IC with that of PMT with known QE.

## Slide 10: Collection Efficiency (Col)
Definition
How to measure
  Measure the Cathode current (IC).
  Add 10-5 ND filter in front of PMT.
  Measure the counting rate of the single PE (S).
  Take the ratio of S1.610-19 105/IC.

## Slide 11: Detective Quantum Efficiency (DQE)
Definition:
    Often confused as QE by “Physicists”
How to measure:
  Use a weak pulsed light source (so that >90% pulse gives the pedestal.)
  Measure the counting rate of the single PE (S).
  Compare S with that of PMT with known DQE.

## Slide 12: Gain (GP)
Definition by Physicists:
How to measure:
  Use a weak pulsed light source (so that >90% pulse gives the pedestal.)
  Measure the center of the mass of Single PE charge distribution of the Anode signal (QA).
  Take the ratio of QA/1.610-19 .
(i = Gain of the i-th dynode)

## Slide 13: Gain (GI)
Definition by Industries:
How to measure:
  Measure the Cathode current (IC).
  Add 10-5 ND filter in front of PMT.
  Measure the Anode current (IA).
  Take the ratio of IA105/IC.
(i = Gain of the i-th dynode)

## Slide 14: Anode Signal (E)
Definition:
(by Industries)
(by Physicists)
(N = No. of Incident Photons)
(Npe = No. of Photo-electrons)

## Slide 15: Energy Resolution (/E)
In ideal case:
In reality:
      QE Quantum Efficiency
      Col Collection Efficiency:
      ENF Excess Noise Factor (from Dynodes)
      ENC Equivalent Noise Charge (Readout Noise)

## Slide 16: Excess Noise Factor (ENF)
Definition:
In case of PMT:
How to measure:
  Set Npe = 10-20 (for nice Gaussian).
  Measure /E of the Gaussian distribution.
  ENF is given by

## Slide 17: Remarks
Don’t try to estimate Npe or N from /E !
(Assuming ENC is negligible.)

## Slide 18: Resolution of Hybrid Photodiode (HPD)
HPD can count 1, 2, 3… PE separately.
  ENF=1.0
But it is still suffering from poor QE.
  We can never beat the Poisson statistics !
200
300
400
500
600
ADC Channel
Pedestal
1 pe
NIM A 442 (2000) 164-170
3 pe
2 pe
4 pe

## Slide 19: Typical Values and Resolution of Various Photon Detectors
 | QE | Col | i | ENF | G | ENC | /E
Ideal | 1.0 | 1.0 | 1000 | 1.0 | 106 | 0 | 1/N
PMT | 0.3 | 0.9 | 10 | 1.3 | 106 | 103 | 4/N
PD | 0.8 | 1.0 | - | 1.0 | 1 | 103 | 1.4/N+(1000/N)2
APD | 0.8 | 1.0 | 2 | 2.0 | 100 | 103 | 3/N+(14/N)2
HPD | 0.3 | 0.95 | 1000 | 1.0 | 103 | 103 | 3/N+(3/N)2
HAPD | 0.3 | 0.95 | 1000 | 1.0 | 105 | 103 | 3/N

## Slide 20: Energy Resolution vs. N
Poisson Limit
Photo Diode
APD
HPD
PMT

## Slide 21: Talk Outline
Introduction: Principle of PMT
NIST Standard Silicon Photodiode
Uncertainties, Concerns Specific to PMTs
Suggestions and Proposal to our Community

## Slide 22: Principle of Silicon Photodiode
Gain = 1.0
QE ~ 100%
Extremely Stable
Large Dynamic Range

## Slide 23: Quantum Efficiencies of NIST Standards (Si, InGaAs and Ge photodiodes)

## Slide 24: Propagation Chain of Absolute Calibration of Photon Detectors
Cryogenic Radiometer
Trap Detector
Pyroelectric Detector
Laser(s)
NIST standard UV Si PD
Reference PMT
Atmospheric Fluorescence
Monochromator(s)
UV LED
Xe Lamp
Laser(s)
Particle Beam
Cosmic-ray Events
PMTs in Telescopes
Light Beam
Scattered Light
Dome Reflector
NIST
US
NIST standard UV Si PD
Standard Light Beam

## Slide 25: NIST High Accuracy Cryogenic Radiometer (HACR)
Shoot a laser to a black body of 4.2oK.
Balance heat by electrical power which produces resistive heating.
Uncertainty is 0.021% (at 1mW).

## Slide 26: Trap Detector
~99.5% efficiency of trapping.
Uncertainty is 0.05%.
Front View
Bottom View

## Slide 27: Scale transfer by substitution method with the HACR

## Slide 28: Propagation Chain of Absolute Calibration of Photon Detectors
Cryogenic Radiometer
Trap Detector
Pyroelectric Detector
Laser(s)
NIST standard UV Si PD
Reference PMT
Atmospheric Fluorescence
Monochromator(s)
UV LED
Xe Lamp
Laser(s)
Particle Beam
Cosmic-ray Events
PMTs in Telescopes
Light Beam
Scattered Light
Dome Reflector
NIST
US
NIST standard UV Si PD
Standard Light Beam

## Slide 29: Spectral output flux of the UV and visible to near-IR monochromators

## Slide 30: UV Working Standard Uncertainty
Transfer from Pyroelectric (relative) and Scaling with Visible WS (absolute).

## Slide 31: Ultraviolet Spectral Comparator Facility (UV SCF)

## Slide 32: Transfer to test (customer) detectors relative combined standard uncertainty

## Slide 33: Example of “Report of Test”

## Slide 34: Absolute Spectral Responsivity of Silicon Photodiode U1xxx
2(%)
S=
=

## Slide 35: Propagation Chain of Absolute Calibration of Photon Detectors
Cryogenic Radiometer
Trap Detector
Pyroelectric Detector
Laser(s)
NIST standard UV Si PD
Reference PMT
Atmospheric Fluorescence
Monochromator(s)
UV LED
Xe Lamp
Laser(s)
Particle Beam
Cosmic-ray Events
PMTs in Telescopes
Light Beam
Scattered Light
Dome Reflector
NIST
US
NIST standard UV Si PD
Standard Light Beam

## Slide 36: Talk Outline
Introduction: Principle of PMT
NIST Standard Silicon Photodiode
Uncertainties, Concerns Specific to PMTs
Suggestions and Proposal to our Community

## Slide 37: Uncertainties Specific to PMTs
PMTs are not perfect. There are many issues to be concerned:
  Cathode and Anode Uniformity
  Wave Length Dependence of QE
  Photon Incident Angles
  Effect of Magnetic Field
  Non Linearity
  Temperature Dependence
  Long-term Stability

## Slide 38: Cathode Uniformity and Area Correction
by HiRes (D.J. Bird et al.)
at =350nm
HiRes PMT
(Photonis XP3062)
NIST SiPD
(UDT UV100)
5mm
4cm

## Slide 39: Sensitivity Map of 16-Pixel PMT for EUSO (R8900-M16F-S12)
by RIKEN/EUSO Focal Surface Subgroup
Total Sum of 16 Pixels
Signal on One Pixel

## Slide 40: Typical QE of HiRes PMTs (Photonis XP3062)
Measured by Photonis
UV LED
N2 Laser
YAG Laser
SKB
YAG Laser
337nm
355nm
370nm
420nm

## Slide 41: Typical Angle Dependence of QE
From Hamamatsu PMT Handbook
Telescope

## Slide 42: Effect of Magnetic Field on Liner-focus 2” PMT
x
y
z
Hamamatsu 2” PMT (R7281-01)
Earth B-Field

## Slide 43: Linearity of ETL 8” PMT
at UCLA PMT Test Facility

## Slide 44: Typical Temperature Coefficients of Anode Sensitivity
From Hamamatsu PMT Handbook
-0.4%/oC

## Slide 45: Typical Long-term Stability
From Hamamatsu PMT Handbook

## Slide 46: PMT vs. Silicon Photo Diode
Source of Systematic Error | Si PD | PMT | Systematic Uncertainty
(Absolute) Quantum Efficiency | ~0.7 | 0.2-0.3 | 3%
Wave-length Dependence of QE | 1% |  10% | 5%
Cathode Uniformity | 1% | 10% | 5%
Photo-Electron Collection Efficiency | 0.99 | 0.8-0.95 | 10%
Gain | 1.0 | 105-7 | 5%
Voltage Dependence of Gain | None | HV6 | 3%
Anode (Gain) Uniformity | 1% | 30% | 10%
Effect of Earth Magnetic Field | None | 10% | 10%
Temperature Dependence | None | -0.4%/oC | 3%
Incident Angle Dependence | 1o | 30o | 10%
Intensity Correction (ND filter) | 1 | 10-5 | 5%
Area Correction | 5mm  | 5cm  | 5%
Non Linearity | None | 5% | 3%
Rate Dependence | None | 5% | 3%
Long Term Stability | Stable | 5%/year | 5%
Total Systematic Uncertainty | 1% |  | 25%

## Slide 47: Talk Outline
Introduction: Principle of PMT
NIST Standard Silicon Photodiode
Uncertainties, Concerns Specific to PMTs
Suggestions and Proposal to our Community

## Slide 48: UCLA PMT Test Facility
15 years of experience to develop and evaluate new PMTs.
  KTeV CsI Calorimeter ¾” & 1½” Linear-focus
  CDF Shower Max Multi-pixel (R5900-M16)
  Auger SD 8-9” Hemispherical
We can measure:
  (Absolute) Quantum Efficiency
  Collection Efficiency
  Gain and Dark Current vs. HV
  Single PE Distribution
  Excess Noise Factor
  Cathode and Anode Uniformity
  Dark Pulse Rate and After Pulse Rate
  Temperature Dependence
  Non Linearity

## Slide 49: Proposal to Evaluate PMTs from HiRes, Auger-FD and EUSO at UCLA
We propose to evaluate the sample PMTs from:
  Auger-FD
  HiRes
  EUSO
  Beam Tests
  Reference PMTs
We plan to add more equipments to measure:
  Effect of Magnetic Field
  Photon Incident-angle Dependence
  Long-term Stability
  …
Mutual Comparison

## Slide 50: How to Minimize Systematic Uncertainties
Don’t try to measure QE, Collection Efficiency and Gain separately.
  Just measure all three together (with a mirror, UV filter etc. as well).  “End-to-end Calibration”
  Prepare the “absolutely-calibrated” light source.
Make sure that the “absolutely-calibrated” light source has the same characteristics as cosmic-ray fluorescence signals.
  Same wave length
  Same angular distribution on the PMT surface
  Uniform over the PMT surface
  Same pulse width and intensity
Calibrate in situ, monitor external environment.
  Same (Earth) magnetic field
  Same temperature
  Same supplied HV

## Slide 51: How to Minimize Systematic Uncertainties at Beam Tests
Don’t try to measure QE, Collection Efficiency and Gain separately.
  Just measure all three together (with a mirror, UV filter etc. as well).  “End-to-end Calibration”
  Prepare the “absolutely-calibrated” light source.
Make sure that the “absolutely-calibrated” light source has the same characteristics as beam test fluorescence signals.
  Same wave length
  Same angular distribution on the PMT surface
  Uniform over the PMT surface
  Same pulse width and intensity
Calibrate in situ, monitor external environment.
  Same (Earth) magnetic field
  Same temperature
  Same supplied HV

## Slide 52: Propagation Chain of Absolute Calibration of Photon Detectors
Cryogenic Radiometer
Trap Detector
Pyroelectric Detector
Laser(s)
NIST standard UV Si PD
Reference PMT
Atmospheric Fluorescence
Monochromator(s)
UV LED
Xe Lamp
Laser(s)
Particle Beam
Cosmic-ray Events
PMTs in Telescopes
Light Beam
Scattered Light
Dome Reflector
NIST
US
NIST standard UV Si PD
Standard Light Beam

## Slide 53: How to absolutely calibrate the PMTs in Telescope
Prepare a stable light source and an “absolutely calibrated photon detector”.
Measure the “absolute photon flux” from the light source, going into the real telescope by the calibrated photon detector above.
Measure the PMT output signal (ADC counts) in the telescope.
Constantly monitor the relative change of the photon intensity from the calibration light source.

## Slide 54: Light Source for Auger-FD
UV LED
Light Source (15cm )
Drum (2.5m , 1,4m deep)
By Jeff Brack

## Slide 55: Drum mounted at the aperture of Bay 4 at Los Leones

## Slide 56: Uniformity of lights emitted from the Drum
By Jeff Brack

## Slide 57: Systematic Uncertainty after End-to-end Calibration
Source of Systematic Error | Si PD | PMT | Sys. Uncertainty Before After | 
(Absolute) Quantum Efficiency | ~0.7 | 0.2-0.3 | 3% | 3%
Wave-length Dependence of QE | 1% |  10% | 5% | 5%
Cathode Uniformity | 1% | 10% | 5% | -
Photo-Electron Collection Efficiency | 0.99 | 0.8-0.95 | 10% | -
Gain | 1.0 | 105-7 | 5% | -
Voltage Dependence of Gain | None | HV6 | 3% | 3%
Anode (Gain) Uniformity | 1% | 30% | 10% | 3%
Effect of Earth Magnetic Field | None | 10% | 10% | -
Temperature Dependence | None | -0.4%/oC | 3% | -
Incident Angle Dependence | 1o | 30o | 10% | 3%
Intensity Correction (ND filter) | 1 | 10-5 | 5% | 5%
Area Correction | 5mm  | 5cm  | 5% | 5%
Non Linearity | None | 5% | 3% | 3%
Rate Dependence | None | 5% | 3% | 3%
Long Term Stability | Stable | 5%/year | 5% | 5%
Total Systematic Uncertainty | 1% |  | 25% | 12%

## Slide 58: Conclusion
Absolute calibration of PMT (in Lab) is absolutely nontrivial. Don’t try it.
The most important is to develop the absolutely calibrated “standard candle” which exactly mimics the physics signal.
Then calibrate your telescope in situ.

## Slide 59: Acknowledgements
Special thanks go to
  Jeff Brack (Pierre Auger FD)
  Stan Thomas (HiRes)
  Hiro Shimizu (EUSO)
  Yoshi Ohno (NIST)
  Yuji Yoshizawa (Hamamatsu)
  and Kai Martens for giving me this opportunity.
This talk is available at
  http://www.physics.ucla.edu/~arisaka/hires/
    pmt_calibration.pdf (2.6M)
    pmt_calibration.ppt (3.3M)