## Slide 1: Vacuum based Photon Detectors
University of California, Los Angeles
Department of Physics and Astronomy
arisaka@physics.ucla.edu
Katsushi Arisaka

## Slide 2: Outline
Concept of Photomultiplier
Basic Properties
  QE, Gain, Time Response
Imperfect Behavior of PMT
  Linearity, Uniformity, Noise…
Other Vacuum Devices
  Hybrid PD/APD
  Applications
Energy Resolution
Summary

## Slide 3: Concept of PMT

## Slide 4: PMT (Photomultiplier Tube)
Notes: PMT is a vacuum photo-sensor, which converts light into electric signal.
The prime features are high sensitivity and fast time response with large area.
Some of Silicon PD have faster time response, but its effective area is very small like 1mm or less.
PMT has variety of sizes, its range is from 10 mm to 500 mm. There are several shapes such as round, square, hexagonal and cylindrical like this.

## Slide 5: Super-Kamiokande
11,200 of 20” PMTs

## Slide 7: Operation of Head-On Type PMT
signal light->photoelectron
photoelectron->Dy1
electron-> multiplication
cascade multiplication
electric signal from anode
Notes: This slide shows basic operation of a HEAD-ON PMT. When signal light enters to a photocathode, it emits photoelectron into vacuum.
The photoelectron is accelerated by electric field inside, and then reaches to first Dynode.
The dynode emits secondary electrons, and they are accelerated to next dynode. This process is repeated like this.
Finally electric signal is taken from anode.
Nobody has seen it yet, but you can see how electrons are moving inside of PMT with this movie.

## Slide 8: Structure of Linear-focus PMT
Mesh Anode
Last Dynode
Photo Cathode
Second Last
Dynode
First Dynode
Glass Window
Photons
QE
CE
1
2
3
n
N
G = 123  n
E=NQECEG

## Slide 9: Principle of Silicon Photodiode
Gain = 1.0
QE ~ 100%
Extremely Stable
Large Dynamic Range

## Slide 10: FAQ
Why still PMT? Why not Silicon Photodiode?
  Intrinsically high gain
  Low noise – photon counting
  Fast speed
  Large area
  but
  Poor Quantum Efficiency
  Bulky
  Expensive

## Slide 11: Purpose of Photon Detector
Observe all the quantities of photons as accurate as possible.
  The number of photons: E
  Arrival time of photons: T
  Position of photons: X, Y, Z
Primary purpose of vacuum detectors:
  Very small number of photons: < 100 photons
  Accurate time of photons: < 10 nsec

## Slide 12: Basic Properties

## Slide 13: Outline
Fundamental Parameters of PMT
  Quantum Efficiency (QE)
  Photoelectron Collection Efficiency (CE)
  Gain (G)
  Excess Noise Factor (ENF)
How to Measure These Parameters
Energy Resolution (/E)

## Slide 14: Quantum Efficiency (QE)

## Slide 15: Quantum Efficiency (QE)
Definition:
The single most important quantity

## Slide 16: QE curves of 6 types
VUV
UV
Visible
Infra-Red
Notes: This slide shows QE curves of 6 types of photocathode. Transverse axis is wavelength and vertical axis is QE.
Bialkali is popular, it’s suitable for scintillation counting. Multialkali (S-20) is also popular, it’s suitable for fluorescence detection.
This is called S-25, that is RED enhanced Multialkali, which is used for NOX detector (air pollution monitor).
Blue one is CsTe, which has solar blind characteristic. These 4 types are called alkali photocathode.
On the contrary, this is GaAsP and this is GaAs. These are called semiconductor photocathode.
GaAsP has highest sensitivity from 400nm to 700nm, but there are some demerits such that large size can’t be made and cost is expensive.
In any case, popular photocathode is bialkali and next is Multialkali photocathode.

## Slide 17: Typical QE
Bialkali:
Sb-Rb-Cs
Sb-K-Cs

## Slide 18: Transmittance of windows
popular
Visible
UV
VUV
More Expensive
Wavelength is Shorter
Notes: This graph shows transmittance of windows, that is faceplates. Transverse axis is wavelength, vertical axis is transmittance.
These are names of wavelength region. VUV (10-200nm) stands for vacuum ultra violet, UV (200-380nm) is ultra violet, and visible (380-780nm) is visible.
The difference of transmittance appears in VUV and UV region.
This Borosilicate glass is popular one. UV and synthetic quartz window are sometimes used.
Depending on the wavelength of signal light, window is selected, but the cost is more expensive when wavelength is shorter.

## Slide 19: FAQ
Why is QE limited to ~40% at best?
  Competing two factors:
    Absorption of photon
    Emission of photo-electrons
  Isotropic emission of photo-electrons.

## Slide 20: FAQ
How can we measure QE?
  Connect all the dynodes and the anode.
  Supply more than +100V for 100% collection efficiency.
  Measure the cathode current (IC).
  Compare IC with that of a reference photon-detector with known QE.

## Slide 21
UCLA QE System
Reference PMT
PMT with unknown QE
Xe Lamp
Source PMT
Monochromator
Integrating Sphere

## Slide 22
UCLA Vacuum UV QE System
PD
PMT
W Lamp
D2 Lamp
Monochromator
UCLA
Hamamatsu

## Slide 23: Propagation Chain of Absolute Calibration of Photon Detectors
Cryogenic Radiometer
Trap Detector
Pyroelectric Detector
Laser(s)
NIST standard UV Si PD
Reference PMT
Real Light Source
Monochromator
UV LED
Xe Lamp
Laser(s)
Particle Beam
Real experiments
PMTs in our detectors
Light Beam
Scattered Light
NIST
us
NIST standard UV Si PD
Standard Light Beam

## Slide 24: NIST High Accuracy Cryogenic Radiometer (HACR)
Photon energy is converted to heat.
Heat is compared with resistive (Ohmic) heating.
0.021% accuracy at 1mW.
This is the origin of absolute photon intensity.

## Slide 25: Trap Detector
Front View
Bottom View

## Slide 26: NIST Standards: Quantum efficiencies of typical Si, InGaAs, and Ge photodiodes

## Slide 27: Sk (Cathode Sensitivity) and Skb (Cathode Blue Sensitivity)
Filter for Skb
Lump for Sk

## Slide 28: Collection Efficiency (CE)
Definition

## Slide 29: FAQ
How can we measure Collection Efficiency?
  Measure the Cathode current (IC).
  Add 10-5 ND filter in front of PMT.
  Measure the counting rate of the single PE (S).
  Take the ratio of S1.610-19 105/IC.

## Slide 30: Detective Quantum Efficiency (DQE)
Definition:
    Often confused as QE by “Physicists”

## Slide 31: FAQ
How can we measure Detective QE?
  Use a weak pulsed light source (so that >90% pulse gives the pedestal.)
  Measure the counting rate of the single PE (S).
  Compare S with that of PMT with known DQE.

## Slide 32: Dynode Structure

## Slide 33: PMT Types
< SIZE >
1/2 inch & 1-1/8 inch
< Features >
Compact
Relatively Cheap
< SIZE >
3/8 inch ~ 20 inch
< Features >
Variety of sizes, Direct coupling
< SIDE-ON TYPE >
< HEAD-ON TYPE >
Notes: There are 2 types of PMTs. The one is SIDE-ON type, and the other is HEAD-ON type.
The name of SIDE-ON comes from that PMT turns ON when light enters from side direction. There are only 2 sizes, ½ inch and 1-1/8 inch.
The features are compactness, relatively cheap price.
The name of HEAD-ON comes from that PMT turns ON when light enters from the head.
One of the features is variety of sizes, its range is from 3/8 inch to 20 inch.
There is photocathode inside of input window, it’s easy to couple a PMT with signal light.

## Slide 34: Dynode Structures – Side-on vs. Head-on
CIRCULAR CAGE
Compact
Fast time response
(mainly for
Side-On PMT)
Good CE
(Good uniformity)
Slow time response
BOX & GRID
< SIDE-ON >
< HEAD-ON >
Notes: There are various dynode structures. Each type has merit and demerit. The dynode structure is selected according to application.
CIRCULAR CAGE : This is most popular dynode for SIDE-ON PMT. Its shape looks like circular cage. Features are compact, fast time response.
BOX & GRID : Its shape looks like a box. As the size is large, it’s easy to collect photoelectron from photocathode.
Therefore, collection efficiency is good. And uniformity is also good. On the contrary, as the size is large, time response is slow.

## Slide 35: Dynode Structures – Linear Focus vs. Venetian Blind
LINEAR FOCUSED (CC+BOX)
Fast time response
Good pulse linearity
VENETIAN BLIND
Large dynode area
Better uniformity
Larger DY1 is used in recent new PMTs (Box & Line)
Notes: LINEAR FOCUSED : This is mixture of CIRCULAR CAGE and BOX & GRID. Its features are fast time response and good pulse linearity.
Larger DY1 is used in recent new PMTs, that is called BOX and LINE dynode.
VENETIAN BLIND : This shape looks like blind at the window in the office. Its features are large dynode area and better uniformity.
This dynode is used for 20 inch PMT because of its large dynode area, but this is not popular.

## Slide 36: Metal Channel PMT
Compact
Fast time response
Position sensitive
PMT with Metal Channel Dynode
16mm in dia.
METAL CHANNEL
Pitch:1mm
TO-8 type PMT
Notes: This is relatively new dynode structure, it’s called metal channel dynode. This dynode was developed 15 years ago as a challenge of Hamamatsu.
As the dynode pitch is 1mm, it can be used for a small size PMT. This is cross-section of TO-8 type PMT.
You can see that there are several layers of dynodes stacked together. Its features are compact, fast time response. As the electron trajectory kept inside of 1 line, position sensitive PMT can be made with this dynode. We have variety of multi-anode series with 1 inch square metal package. Our Flat panel PMT is using second generation of metal channel dynode.

## Slide 37: Fine Mesh PMT
Fine Mesh

## Slide 38: MCP (Micro Channel Plate)
Gain = 100 - 1000
( 5 – 10 μm ϕ)
MCP

## Slide 39: MCP PMT
MCP PMT
Image Intensifier

## Slide 40: Principle of Image Intensifier
http://www.e-radiography.net/radtech/i/intensifiers.pdf

## Slide 41: Effect of Magnetic Fields
Metal
Channel
Fine
Mesh
MCP
PMT
Solid
State
Linear
Focus
HPD
APD

## Slide 42: Gain of PMT

## Slide 43: Structure of Linear-focus PMT
Mesh Anode
Last Dynode
Photo Cathode
Second Last
Dynode
First Dynode
Glass Window
Photons
QE
CE
1
2
3
n
N
G = 123  n
E=NQECEG

## Slide 44: Secondary electron Emission
  HV0.6

## Slide 45: Gain (GP)
Definition by Physicists:
(i = Gain of the i-th dynode)

## Slide 46: FAQ
How can we measure the Gain (GP) of our definition?
  Use a weak pulsed light source (so that >90% pulse gives the pedestal.)
  Measure the center of the mass of Single PE charge distribution of the Anode signal (QA).
  Take the ratio of QA/1.610-19 .

## Slide 47: Single PE distribution

## Slide 48: Gain (GI)
Definition by Industries:
(i = Gain of the i-th dynode)

## Slide 49: FAQ
How do manufactures measure the real Gain (GI)?
  Measure the Cathode current (IC).
  Add 10-5 ND filter in front of PMT.
  Measure the Anode current (IA).
  Take the ratio of IA105/IC.

## Slide 50: Gain vs. Voltage Curve
Physicists Definition:
GP=δ1•δ2•… •δn
Industries Definition:
GI=CE•δ1•δ2•… •δn
CE=GI/GP~80%.
GP by UCLA
GI by Photonis

## Slide 51: 270 Auger-SD PMTs: HV for G=2105 UCLA vs. Photonis
HV varies from PMT to PMT.
Photonis is Higher than UCLA (due to CE).
CE varies from PMT to PMT.
UCLA
Photonis

## Slide 52: FAQ
Why is the Gain so different from PMT to PMT at the fixed HV?
  At given HV, each  may be 10% different.
  Then, Gain could be an order of magnitude different. (G = 123  n)

## Slide 53: FAQ
What is the maximum allowed HV for stable PMT operation?
  It can be checked by Dark Current behavior.

## Slide 54: Gain and Dark Current vs. HV
Thermal
Photoelectron
Emission
Leakage
Current
Field
Effect

## Slide 55: Temperature Dependence of Anode Sensitivity
-0.4%/oC

## Slide 56: Two Types of Voltage Divider
<-HV Operation>
<+HV Operation>
Pulse operation only
No DC output
Notes: There are 2 types of voltage dividers. The one is for –HV (negative HV) operation, the other is for +HV (positive HV) operation.
As electron can be accelerated by plus voltage, positive voltage is always applied to anode.
The matter is which side is connected to ground. If anode side is connected to ground, cathode side has to be negative voltage.
If cathode side is connected to ground, anode side has to be positive voltage.
In case of –HV, as anode side is connected to ground, it’s easy to read out signal because there is no potential difference between anode and read out circuit.
In case of +HV, as cathode side is connected to ground, it’s easy to hold a PMT because there is no potential difference between cathode and holder. It results in low noise.
On the contrary, there is potential difference between anode and read out circuit. A coupling capacitor has to be placed between them.
As a results, only pulse output can be obtained. (no DC output) When count rate is high, baseline shift might be occurred.

## Slide 57: Time Response

## Slide 58: Time Response
TTS
Transit Time Spread
(Variation of Transit Time)
Transit Time
RISE TIME
10% to 90%
FALL TIME
90% to 10%
Example of
Waveform
Rise : 1.5 ns
Fall : 2.7 ns
Notes: One of PMT prime features is fast time response. This slide shows DELTA FUNCTION LIGHT and its PMT OUTPUT.
There are some definitions for PMT time response.
When pulsed signal light with very narrow width is enters to PMT, it generates output like this.
Rise time is defined as the time interval between 10% to 90% of PMT output.
Fall time is defined as the time interval between 90% to 10%.
Transit time is defined as the time interval between DELTA FUNCTION LIGHT and peak of PMT output.
TTS, it stands for Transit Time Spread, it’s variation of transit time. TTS is defined as the FWHM of TT variation.
In case that a scintillator with fast time response is used or count rate is high, fast time response is required.

## Slide 59: Typical TTS (Transit Time Spread)

## Slide 60: Transit Time vs. HV
Higher Voltage

Faster Transit Time

## Slide 61: Time Properties (R11410)

## Slide 62: Time Resolution vs. Sensitive Area
HPD
SiPM

## Slide 63: Imperfect Behavior of PMT

## Slide 64: Uncertainties Specific to PMTs
PMTs are not perfect. There are many issues to be concerned:
  Non Linearity
  Cathode and Anode Uniformity
  Effect of Magnetic Field
  Temperature Dependence
  Dark Counts
  After Pulse
  Rate Dependence
  Long-term Stability

## Slide 65: Linearity

## Slide 66: PMT Non Linearity
Non Linearity is the effect of the space charge mainly between the last and the second last dynode.
Mesh Anode
Last Dynode
Photo Cathode
Second Last
Dynode
First Dynode
Glass Window
Photons
QE
Col
1
2
3
n
N

## Slide 67: Pulse Linearity
What is Pulse Linearity ?
Relation between radiation energy and PMT output.
Deviation from ideal line (%)
PMT output / peak current (mA)
PMT output
Radiation Energy
Light Intensity
Notes: What is Pulse Linearity ?? It’s relation between radiation energy and PMT output.
In case that PMT is used for scintillation counting and spectrum of radiation energy is measured, pulse linearity is important characteristic.
PMT output becomes large when radiation energy becomes large. If pulse linearity is not so good, PMT output doesn’t show correct radiation energy.
This graph shows deviation of actual PMT output compared with expected (or ideal) PMT output.

## Slide 68
Block Diagram for Double-Pulsed Mode
Dim: 1
Bright: 4
Pulse Linearity
Notes: This is block diagram for Double-Pulsed Mode to measure pulse linearity. Bright and dim pulsed light sources are used.
When PMT output is small, the ratio between bright and dim is ideal. It’s usually the ratio of 4 to 1.
When PMT output is getting larger while changing the distance between light sources and a PMT, the bright-dim ratio becomes larger like 4.2 or smaller like 3.8.
This means that pulse linearity becomes worse. Pulse linearity depends on type of dynode structure. Linear focus type is good in this view point.

## Slide 69: Optimization of Anode Pulse Linearity
(The last 3 stages)

## Slide 70
Linearity at different gains
Low gain (1000V)
High gain (1500V)

## Slide 71: Uniformity

## Slide 72: Anode Uniformity
spot light
SLIT shape
Incident light
Large size of
Incident light
Notes: This is anode uniformity. Output current from anode is measured while spot light is scanned on photocathode.
HEAD-ON PMT has relatively good uniformity, but SIDE-ON PMT doesn’t have it. This is due to the difference of PMT construction.
SIDE-ON PMT is usually used for spectrometer, and the size of signal light is small in that application.
On the contrary, HEAD-ON PMT is usually used for scintillation counting, the size of signal light (that is size of scintillator) is relatively large.
Therefore, anode uniformity is important characteristics for HEAD-ON PMT.
Fortunately, there is space for focusing in HEAD-ON PMT, we can apply suitable design to improve the uniformity.

## Slide 73: Cathode Uniformity (3 inch PMT)

## Slide 74: Anode Uniformity (3 inch PMT)

## Slide 75: Collection Efficiency (=Anode/Cathode)
(KA0044)

## Slide 76: Effect of Magnetic Field

## Slide 77: Effect of Magnetic Fields
Metal
Channel
Fine
Mesh
MCP
PMT
Solid
State
Linear
Focus
HPD
APD

## Slide 78: Typical Magnetic Field Effect
Earth B-Field

## Slide 79: Effect of Magnetic Field on Liner-focus 2” PMT
x
y
z
Hamamatsu 2” PMT (R7281-01)
Earth B-Field

## Slide 80: Edge Effect of Magnetic Shields
For effective shielding,
we need extra mu-metal
in front.

## Slide 81: Dark Count

## Slide 82: Temperature Dependence of Dark Current

## Slide 83: Dark Count Rate vs. Temperature

## Slide 84: After Pulse

## Slide 85: After Pulse (R11410)

## Slide 86: After Pulse by Helium
Helium Contaminated PMT from MACRO
> 10%

## Slide 87: Long Term Stability

## Slide 88: Typical Long-term Stability
From Hamamatsu PMT Handbook

## Slide 89: Other Vacuum Devices

## Slide 90: Principle of Silicon Photodiode
Gain = 1.0
QE ~ 100%
Extremely Stable
Large Dynamic Range

## Slide 91: APD (Avalanche Photodiode)
High Gain (100-1,000), High QE (~70%).
Then, why not replace PMTs?
Drawbacks:
   <2, ENF>2  Effectively QE <35%.
  Extremely Sensitive to Temperature and Voltage change.
  Difficult to manufacture uniform, large area.

## Slide 92: HPD (Hybrid Photodiode)
In vacuum, Silicon Photodiode instead of dynodes.
High Gain (1000-3000), we can count 1-5 photoelectrons.
Then, why not replace PMTs?
Photo Cathode
Silicon PD
e-

## Slide 93: CMS Detector under 4 Tesla
4 Tesla
EM
Hadron
APD
HPD

## Slide 94
CMS HCAL Multi pixel HPD (DEP)
PIN Diode
array
Ceramic feedthrough
Fiber-Optic
Window
Photocathode (-10 kV)
e
19 channel pixel layout
pixel size: 5.4 mm flat-flat
gap between pixels: 0.04 mm
3.4 mm

## Slide 95: LHCb experiment
RICH
RICH

## Slide 96
The pixel HPD by DEP (for LHCb)
Advantages of this hybrid, pixel structure:
low noise: excellent resolution of single photoelectrons
high channel number/density
DEP, The Netherlands

## Slide 97: Hamamatsu Hybrid APD
Single Channel HAPD
64 Channel HAPD
+ Readout

## Slide 98: 1, 2, 3 … Photo-electron Distribution
1
2
3
4
5
6 Photo-electrons
True photon counting

## Slide 99: Decay Time Measurement by HAPD
Time Resolution = 80 psec
FWHM = 1.5 ns
No after pulse
Pulse Shape
Decay Time

## Slide 100: Leica HyD Detector for Confocal Microscope
Hamamatsu
Compact HAPD
with GaAsP

## Slide 101: 8 inch HAPD by Hamamatsu
New release
at NSS 2012

## Slide 102
Water Tank
Liquid
Scinti
Water Tank
Xe 20 ton (10 ton)
40Ar 70 ton (50 ton)
15 m
6 m
Liquid
Scinti
Xe
Ar
MAX G3 Dark Matter Detector

## Slide 103: QUPID (QUartz Photon Intensifying Detector)
Photo Cathode
(-6 kV)
APD (0 V)
Quartz
Quartz
Al coating
APD (0 V)
Photo Cathode
(-6 kV)
Made by Synthetic Silica only.

## Slide 104: Production Version QUPID

## Slide 105: 1, 2 and 3 PE Distribution with 2m cable
2 PE
3 PE
1 PE
G = 800 × 200
= 160,000
TTS = 160 ps (FWHM)

## Slide 106: Intevac Electron Bombarded CMOS

## Slide 107: EBAPS by Intevac

## Slide 108: Energy Resolution

## Slide 109: Anode Signal (E)
Definition:
(by Industries)
(by Physicists)
(N = No. of Incident Photons)
(Npe = No. of Photo-electrons)

## Slide 110: Energy Resolution (/E)
In ideal case:
In reality:
      N Number of incident photons
      QE Quantum Efficiency
      CE Collection Efficiency:
      ENF Excess Noise Factor (from Dynodes)
      ENC Equivalent Noise Charge (Readout Noise)
      G Gain

## Slide 111: Excess Noise Factor (ENF)
Definition:
In case of PMT:
How to measure:
  Set Npe = 10-20 (for nice Gaussian).
  Measure /E of the Gaussian distribution.
  ENF is given by

## Slide 112: Single PE Distribution
 of the single PE distribution is given by
Thus ENF is related to Peak to Valley Ratio.

## Slide 113: Single PE Distribution
To see single PE, tune light intensity so that >90% gives pedestal.
If 1 >>5, ENF<1.4, Clear single PE can be seen.
The true position is given by the “Center of Mass” including signal below the pedestal.

## Slide 114: ENF vs. P/V Ratio of 270 Auger-SD PMTs

## Slide 115: FAQ
When should we use PMT, and when should we use Silicon Photodiode?
  Depends on intensity of photons.
  Depends on speed of signals.

## Slide 116: Resolution of Hybrid Photodiode (HPD)
HPD can count 1, 2, 3… PE separately.
  1 >1000, ENF=1.0
But it is still suffering from poor QE.
  We can never beat the Poisson statistics !
200
300
400
500
600
ADC Channel
Pedestal
1 pe
NIM A 442 (2000) 164-170
3 pe
2 pe
4 pe

## Slide 117: Summary Table
 | QE | CE | i | ENF | G | ENC | /E
Ideal | 1.0 | 1.0 | 1000 | 1.0 | 106 | 0 | 1/N
PMT | 0.35 | 0.9 | 10 | 1.3 | 106 | 200 | 3.8/N
PD | 0.7 | 1.0 | - | 1.0 | 1 | 200 | 1.4/N+(280/N)2
APD | 0.7 | 1.0 | 2 | 2.0 | 100 | 200 | 2.9/N+(2.9/N)2
HPD | 0.5 | 0.9 | 1000 | 1.0 | 103 | 200 | 2.2/N+(0.4/N)2
HAPD | 0.5 | 0.9 | 1000 | 1.0 | 105 | 200 | 2.2/N
SiPM | 0.7 | 0.4 | 1000 | 1.3 | 106 | 1000 | 4.3/N
VLPC | 0.7 | 1.0 | 1000 | 1.0 | 105 | 200 | 1.4/N

## Slide 118: Energy Resolution vs. N
Poisson Limit
Photo Diode
APD
HPD
PMT

## Slide 119: Resolution (over Poisson Limit)
PMT (35% QE)
HPD (50% QE)
APD
PD
HAPD
VLPC
SiPM
G-APD

## Slide 120: Summary

## Slide 121: Purpose of Photon Detector
Observe all the quantities of photons as accurate as possible.
  The number of photons: E
  Arrival time of photons: T
  Position of photons: X, Y, Z
Primary purpose of vacuum detectors:
  Very small number of photons: < 100 photons
  Accurate time of photons: < 10 nsec

## Slide 122: Market Price
SiPM
Silicon
HPD

## Slide 123: FAQ’s
Why do we have to operate each PMT at different HV?
Why is PMT response non-uniform over surface?
What is the cause of non-linearity?
How stable is PMT? How often should we calibrate? Every minute? Every day??
What external facts could change the Gain of PMT?
What could damage PMTs permanently?

## Slide 124: More FAQ’s
What is the source of dark current and dark pulse?
Are they correlated?
Why is PMT still the best for photon counting application?
Why is APD or HPD not widely used?
Then, who uses APD or HPD?
Why is the signal of PMT so fast?

## Slide 125: Closing Remarks
PMTs are still used in many applications for good reasons:
  Intrinsically high gain
  Extremely low noise – photon counting
  Fast speed ( < 1 ns)
  Large area ( >> 5 inch)
However PMTs are not perfect. There are many issues to be concerned:
  Cathode and Anode Uniformity
  Non Linearity
  Effect of Magnetic Field
  Long-term Stability

## Slide 126: References
Hamamatsu PMT Handbook
  http://sales.hamamatsu.com/assets/applications/ETD/pmt_handbook_complete.pdf
Special thanks to
  Yuji Yoshizawa at Hamamatsu Photonics