5-Nov-03
Version 0.0

Title:

Trigger Studies for the CMS Cathode Strip Chamber Muon System

Likely Authors:

UCLA: Boyd, Cousins, Gilmore, Haapanen, Hauser, Mey, Mohr, Mumford, Valouev

Rice: Matveev, Padley, Roberts, Veramendi, ?

Florida: Acosta, Drozdetski, Scurlock, Stoeck?

Ohio State: Bylsma, Durkin, Gilmore, Gu, Ling, ?

Abstract:

The performance of the Compact Muon Solenoid (CMS) cathode strip chamber (CSC) trigger electronics was tested in the summer of 2003 in a test beam at CERN. A novel feature of the muon and pion beams was a 25 ns bunch structure similar to that of the LHC. Two CSCs were equipped with production on-chamber electronics and connected to near-final prototype off-chamber electronics. Both inputs and outputs of the CSC trigger electronics were recorded. The performance of the trigger electronics is presented as a function of chamber angles, particle rates, chamber high voltage, thresholds, and other variables. Also, the CSC trigger output data is compared bit-for-bit against a detailed simulation of the digital electronics. These studies show a number of interesting effects, and verify that the design of the system is sound and will yield efficient muon triggering in CMS operation at the Large Hadron Collider.

Introduction:

The CSC muon system of CMS (refs – Muon TDR etc) and the plan for reading out and triggering with this system (refs – Trigger TDR etc) have been described in detail elsewhere. The on-chamber electronics for the CSC muon system of CMS were previously studied extensively in (1999, 2001, refs?), and are now being mass-produced and installed on the CSC chambers. However, the performance of the associated off-chamber electronics needs to be extensively checked before their mass production begins. The TMB, DMB, and MPC modules are in the near-final prototype stage, and a test beam with LHC-like time structure is highly desirable for assuring that they will operate well under LHC conditions.

In particular, readout and triggering of this 468-chamber system will require 60 VME 9U-size electronics crates mounted around the periphery of the endcap muon iron disks. For each CSC chamber, the on-chamber electronics is connected to one pair of boards in a peripheral crate: a Trigger MotherBoard (TMB) module and a Data acquisition MotherBoard (DMB) module. A schematic of the CSC electronics system is shown below in Figure 1. 

[image: image1.jpg]Cathode LCT - Trig Motherboard = ------- 1 DAQ Motherboard

Clock Control Board ----

-- PC Interface Board

) (DAQ datahP'CI Bui .

Slow Control

Muon Sector Receiver «
Lev-1 Trigger

Trigger -Timing -Control
(TTCrx)

", '\ME Crate

Cathode Front -end Board
3x6 ASICs, 2 FPGAs,
channel Links, Buffers

Anode LCT Board
FPGA’s, ASICs

Anode Front -end Board
1 ASIC





Figure 1. A block diagram of the electronics associated with each CSC muon chamber of CMS.

Therefore, the key trigger goals for the summer 2003 test beam studies were to show that the CSC peripheral crate electronics: 

1. Work well together as a system.

2. Trigger on muons with high efficiency.

3. Correctly identify the LHC bunch crossing with high probability.

4. Can handle the maximum particle rates expected at LHC. 

A short explanation of the function of each of the modules that are shown in Figure 1 and that were used in the test beam studies follows: 

· CFEB (Cathode Front-End Board):  Contains sensitive amplifiers for cathode strip signals and creates parallel, independent data and trigger paths.  For the precision data path, analog charge information is stored in a switched capacitor array and then digitized for readout. The digitized charge data are then sent to the DMB.  For the trigger data path, custom comparator ASICs find the muon position on each CSC layer to a precision of one half-strip by comparing cathode signals on adjacent strips. The half-strip bits are then sent to the TMB board. There are up to 5 CFEBs per CSC chamber.

· AFEB (Anode Front-End Board): Contains a single amplifier plus constant-fraction discriminator ASIC to digitize anode information. The hits are sent to the ALCT board. There are up to 42 AFEBs per CSC chamber.

· ALCT (Anode Local Charged Track):  For the trigger, this module finds patterns among the six layers of anode hits that look like a muon stub, and not background neutron-induced or other types of hits. It also time-aligns the anode hit information with the LHC clock, and determines muon bunch crossing using a multiple-layer coincidence timing technique. Position and timing information for up to two anode LCT hit patterns (also called ALCTs) are sent to the TMB. For the data readout, the anode hits and the hit patterns found are sent upon receipt of a level 1 trigger decision (L1A) to the TMB.

·   TMB (Trigger Mother Board):  For the trigger, the TMB finds up to two cathode patterns, then requires a time coincidence between anode and cathode patterns.  If a coincidence is found TMB combines the trigger information and sends the best two LCTs to the MPC using the more precise anode bunch crossing time. For the data readout, the TMB passes the anode ALCT information directly to the DMB, and sends in parallel the cathode comparator hits, cathode trigger patterns, and anode-cathode coincidence information to the DMB.

· MPC (Muon Port Card):  Collects LCTs from each of up to nine TMBs in a trigger sector and chooses the best three based on the muon stub quality.  Sends this information to the CSC track finder system (Sector Processor, Sector Receiver).

· CCB (Clock and Control Board):  Provides the interface of the CSC system with the CMS Trigger, Timing and Control (TTC) system.  Distributes necessary signals for synchronized operation of a peripheral crate.

· DMB (Data acquisition Mother Board):  Upon arrival of L1A, collects data from ALCT, TMB and CFEBs, containing ALCTs, CLCTs and analog cathode information from a single CSC.  Sends  this event information to the DDU.

· DDU (Detector-Dependant Unit):  Upon arrival of L1A, collects data from all DMBs in a CSC sector and sends the information down global DAQ path. In the version present at the test beam, this module was read out via Gigabit Ethernet to a PCI card and from there to disk on a Linux computer.

Test Beam Setup:

Two CSCs were equipped with production on-chamber electronics and connected to near-final prototype off-chamber electronics in the X5A test beam, which is a tertiary beam from CERN’s SPS (400 GeV/c), providing a muon or pion beam of energy between 5 and 250 GeV. Collimators in the beam line allowed for control of the rate of particles. A novel feature of the muon and pion during part of the running time was a 25 ns bunch structure similar to that of the LHC. The bunches were 2.3 ns wide, and 48 bunches were filled out of the SPS total of 924 (the LHC has 3564 bunches). Particles were extracted during a 1.5-2.5 s spill out of a 16.8 s ramp cycle. Muon rates up to ~104 per spill and pion rates up to ~106 per spill were available.

The test beam setup is shown in Figure 2. The two CSCs were placed about a meter apart and nominally rotated 20o with respect to the perpendicular to the beam axis, and vertically oriented, so that the beam represented an LHC muon at 20o polar angle and infinite momentum. All parts of the diagram shown in Figure 1 were implemented, including the Muon Port Card (MPC) that collects trigger information about muon segments from all of the TMB modules in a peripheral crate.

The trigger electronics was set to form triggers from internal chamber information, but the readout was initiated by a three-fold coincidence of signals from the scintillator paddles of the beam hodoscope. The size of these paddles was approximately 10 cm on each side, well-matched to the size of the beam. The background rate of non-particle coincidences from this hodoscope was so low as to be unmeasured. The operation of the trigger electronics is checked by study of the data read out. In the case that for some reason there was no CSC data, there was nonetheless an empty header, so that true efficiencies could be measured. 

[image: image2.jpg]



Figure 2. The summer 2003 CSC test beam setup. The beam line is vertical and passes through two CSC chambers that are mounted on the left (one is hidden behind the other). Blue cables run along the top from the on-chamber electronics to the right side where the peripheral crate electronics is located in a 9U-size VME crate.  The peripheral crate electronics is then read out through an optical fiber to a PCI card, and thence to disk in the data acquisition computer in the control room.

CSC event data is organized into three sub-categories: trigger anode data, trigger cathode plus anode/cathode time-coincidence data, and cathode precision charge-readout data. For each of these categories, an 8- or 16-bin time history of all channels is recorded. For anode and cathode trigger data, the time bins are one (25 ns) clock wide, while for the cathode trigger data, the time bins are two (50 ns) clocks wide. A representative event display with a single muon track is shown in Figure 5. The precision cathode readout data develops with the approximately 150 ns rise time of the sensitive front-end amplifiers, while the cathode (and anode???) trigger data is shown as individual hits at the time at which the signals exceeded threshold. The time and position of the muon stub that was found by the trigger electronics is shown in blue (???).

[image: image3.png]ALCT bits vs. Positign ALCT bits vs. Positign





Figure 3. Anode data from a muon event: chamber 1 on left, chamber 2 data on right. The six layers of the chamber are shown, with time bin running towards upper left, wire number towards right.

[image: image4.jpg]CFEB vs. Time

CFEB vs. Position i

CFEB vs. Time

| CFEB vs. Position i

Livaslanaal MITTRI NN sl ok 1 [FTETAIRETACRRTIITT Livaalosnaleassl Lisaal TN 1 1 Lisaalesnalens
% § & & 8 - § & & &8 8 ° § § & B 8 -° T g8 & & g 8 & ¢@ T § & &8 =8 -
- 3 TR
B B 1 - 1
B B 1B

+*
o
Lovaalavaalnesl
§ § &8 8
+ - + -
* +* * *

10

onlinbied balols NP polinlinde
T & 8 T % g & 8 g s
1 "
+ L] + 2 B
L
] - 5
3 e 2 .T o
e
F B L
PP AP PPP AT PO Ll P TP TS PR O O T PP |
§ 8§ ¥ E s T 8 : § &8 8 &8 = § & ¢





Figure 4. A test beam event display with space and time profiles of cathode data. Chamber 1 data on left, chamber 2 data on right. Projected data for the six chambers are stacked vertically, with strip projections on left and time bin projections on right for each chamber.

[image: image5.png]Sswtd Cap,CLCT bitsBest Muon vsksitio Swtd Cap,CLCT bitsBest Muon vssitio

ALCT rav hits ALCT rav hits





Figure 5. A test beam event display showing cathode hit (top) and anode hit (bottom) profiles. The time development of the signals is not shown in this projection. The cathode trigger bits are also shown…why not anode… maybe don’t want this one?

Figure 6. A test beam event display showing cathode time and space development…THESE FILES WERE TOO BIG AND PRACTICALLY IMPOSSIBLE TO CONVERT TO REASONABLE GRAPHICS FORMAT FOR WORD.

Trigger Setup and Timing:

The first step in taking data is to time in the electronics system. A number of times need adjustment (this can be done either with cosmic ray hodoscope or test beam hodoscope triggers):

· CFEB readout timing: a CLCT data-available signal from the TMB is sent through the DMB to the CFEBs. This signal freezes data in the CFEB switched capacitor arrays until a subsequent L1A arrives exactly 2.9 s ±1 bunch crossing later. At the test beam, this is monitored using test points on the CFEB boards and adjusted by changing the CCB L1A delay setting. This adjustment requires an operating CSC chamber and an external (e.g. scintillator) muon trigger.

· ALCT-TMB data transmission timing: data is multiplexed to 80 MHz in both directions between these boards. Two clock phases need adjustment for proper data transfer. The TMB supplies one main (transmit) clock for the ALCT board. To determine the proper phase of this clock, it is delayed in 1 ns steps using a PHOS4 delay ASIC chip (ref.) until ALCT output data is correctly latched at the TMB. In addition, some data is transferred from TMB to ALCT during normal operation, and in order that this data is received correctly at the ALCT a second (receive) clock phase is set using another PHOS4 delay chip channel. The TMB is read out through VME. This timing can be correctly adjusted without particles.

· CFEB comparator chip clock timing: the comparator chip clock on the CFEB is sent from the TMB. The phase of this clock determines whether the data is properly latched at the TMB. Fine timing to get optimal phase of this clock is adjusted by a scan across 25 PHOS4 delay settings. The quality of the timing is monitored by reading out the TMB through VME and counting the rate of 5-layer and 6-layer CLCT patterns. (A working CSC chamber and cosmic rays or test beam events are needed for this step).

· ALCT L1A delay setting: Once the ALCT is properly timed in, the ALCT L1A delay can be set.  Events are taken at different L1A delay settings and the peak in ALCT efficiency is determined.

· TMB: L1A timed in (look at L1A window using internal logic scope and center by adjusting the window). Look at the “l1a_pulse_dsp” signal and make sure it is in the “l1a_window_dsp.”   Adjust the L1A pulse delay in VME address 74.  Also look at the “alct_vpf_tp” pulse and make sure it is in the middle of the “clct_window_tp.”  Adjust the ALCT valid pattern pulse delay in VME address B2.
· Additional DMB timing: active FEB flags for ALCT and CLCT have to be timed in – direct signals on pins on DMB utility board (ugh!).

· Fine timing adjustment of ALCT delays:

[image: image6.png]ALCT Delay Chip Scan

Chamber 1 -- ALCT delay scan Chamber 2 -- ALCT delay scan
- 100 - 100
£ &
o] ol
§ % § %
g g
80 80
T0 T0
60 60
50 S0
% 5 0 5 20 25 30 35 40 30 35
Time Delay (ns) Time Delay (ns)




Figure 7. Anode trigger efficiency versus fine delay setting (ns).

Checks of BX crossings in various places?

Data-taking conditions:

Results:

Efficiency

HV scan

Comparator and anode threshold scans

Angle scans

After timing in, the plot below shows that there were 48 filled bunches out of the 924 bx orbit:

[image: image7.png]Synchronized Beam Bunch Structure

Chamber 1 -- Full CLCT Bxn

e
[&]
-l 2
Q10
10
1

U] 100 200 300 400 500 600 70O 800 900
TME LCT bxn

Chamber 1 -- Zoom CLCT Bxn

CLCTS

! 600 620 640 6G60 680 700 720 740
TME LCT bxn

Chamber 2 -- Full CLCT Bxn

g
=
<10
10
1

0 100 200 300 400 500 600 700 800 900
TME LCT bxn

Chamber 2 -- Zoom CLCT Bxn

CLCTS

1 600 620 640 660 680 700 720 740
TME LCT bxn




Figure 8. The time of arrival of the muon, as seen by the cathode trigger, in clock ticks with respect to the “LHC bunch 0”.

[image: image13.png]HalfStrip vs HalfStrip |

g

8

Chamber 2

]

g
T

8 & 8 8

I R EEYE

L L I
100 120 140 160
Chamber 1

65

Chamber 2
T

55

a L L L L
00 105 110 115 120 125
Chamber 1

Mon Jun 2 20:36:11 2003



Then BX efficiency:

[image: image14.png]Timing |

[ om0 Count-rorrg Bt Gt ] et [ LA X ot ProTg 530 oumn ] (o

W L P R T T e 00
e 1es| £ en 1102
s o2 s ossn

w0l E

10'L E|

10112 114 116 118 120 122 124 126 128 130

110112 114 116 118 120 122 124 126 128 130

BXnumber BXnumber
“ALCT Full BXN -5 BUBXN “ALCT Full BXN -5 BUBXN
| Chamber 1 | Chamber2
g T T T T T Entes 11134 | 8 T T T T T Entries 11140
o
: wen 0o £ O weon o5t
ans 010 ans o1z
w0l E|
10'L E|
w0} E|
ol E
TN I T TN TV [l TP TP | T T TN :
™ o 10 ™ @ > 0 o o m @
BXnumber BXnumber

0:40 2003



Figure 9. Difference in time of the arrival of the muon (in 25 ns clock ticks) between the anode (ALCT) and cathode (CLCT) trigger circuitry and the timing from the scintillator hodoscope.

Figure 10. Position of chamber 1 with respect to chamber 2:

Plots from Erin:

[image: image8.png]File

State || Page || Landscape [ 1.000] | Letter | fCanvasps: HY Scan If Fri Aug 22 16:10:02 2003

pen

Print Markee!

Save Al

Vriabe Size

193k

Save Marked!

ALCT Efficiency vs.Comp.Threshold when Voltage 3600V I
Chamber 2

2
g
3
g

g g
e 7 100,
£ £
H H
oo = soal
g g
H H
ol 996/
4l
02
992l
o)
ool
8sl
o8l
WO i i I i | I L 1 I I
W ow ™ 0 w0 & W o®w w0 w0 &

Camp Threshold (mV) Comp Threshold (V)





Figure 11. Graph of ALCT efficiencies vs. Comparator Threshold values while high voltage is constant at 3600 V in Chamber 1 and    Chamber 2 respectively.

[image: image9.png]File

Vriabe Size

—

apen
Print AL

Print Merked
Save AL

Save Merked

T

Recioplay

1.000/ | Letter fCanvas ps: HY Scan

Wed Aug 6 11:59:38 2003

CLCT Efficiency vs.Comp.Threshold when Voltage 3600V ]

CLT Efficiency (%)

%5 L5
2
7
]
ol 2 oal
°
&
3
a7s[ a7s
ol a7l
5[ 5[
ol 96|
P O PRI PO POV DN P I P PO POV DN
RN N

Comp.Threshod (mV) CompThreshold (mV)





Figure 12. Graph of CLCT efficiencies vs. Comparator Threshold values while high voltage is constant at 3600 V in Chamber 1 and Chamber 2 respectively.

[image: image10.png]3 ey: /home/guest/Erin/FinalALCTeff. ps HV Scan,

File | |state|| Page||Landscape| | 1.000] | Letter /nome/guesYENN/FnalALCTeflps HY Scan Fri Aug 1 15:14:17 2003

Vriabe Size

107k 128

apen
Print AL

Print Merked
Save AL

Save Merked

\_+ ALCT Efficiency vs. High Voltage 1
Chamber 1 Chamber2

101 101
£ £
& 1l & 1o0a-
5 100t 5 1006+
g g
= tmaf = sl
1002 102
00|
90
903
90.4)
90
ol I 1 i I ol I 1 i I
sm  ms  asw  sw s W ms asw  ss0  amo

High Voltage (v) High Volage (V)





Figure 13. Graph of ALCT efficiencies vs. High Voltage values for Comparator Threshold values 30mV, 40mV, and 50mV in Chamber 1 and Chamber 2 respectively.
[image: image11.jpg]CLCT Quality

CLCT Quality(%)

CLCT Pattern(%)

100
90 F
80 CLCT Qualities
70 — halfstrip
60 - - - distrip
- -e- Quality 0
: -o- Quality 1
40 ~e- Quality 2
30 —o- Quality 3
20 E —— Total
10
0
CLCT Pattern
100
90 |
80k CLCT Patterns
70E — halfstrip
; - - - distrip
60 —o— Pattern 1
: —eo— Pattern 2
50 —e— Pattern 3
: —e— Pattern 4
40 —o— Pattern 5
30F —o— Pattern 6
—o— Pattern 7
20 F —— Total

10





Figure 14. CLCT pattern usage and quality versus phi_b for theta=20o from Silja.

[image: image12.jpg]CLCT Quality

100

20

80

CLCT Quality(%)

70

60

50

40

30

20

10

ot

CLCT Pattern

100

920

80
70

CLCT Pattern(%)

60

50

40
30

Phi

CLCT Qualities
— halfstrip

- - - distrip

-o- Quality 0
-o- Quality 1
- Quality 2
- Quality 3
—=— Total

CLCT Patterns
— halfstrip
- - - distrip
—eo— Pattern 1
—o— Pattern 2
—o— Pattern 3
—o— Pattern 4
—e— Pattern 5
—e— Pattern 6
—e— Pattern 7
—=— Total





Figure 15. CLCT pattern usage and quality versus phi_b for theta=40o from Silja.

Conclusions:

References:

BELOW IS MISC TEXT NOT INCORPORATED INTO THE PAPER YET.

Trigger debugging with individual chamber operations: 

· Timing-in procedure is well-defined and debugged. Among the procedures: 

· CFEB-TMB data transmission: run PHOS4 delay curves to make sure comparator bits are latched at exactly the right phase (optimizing 6-hit halfstrip track fraction). 

· ALCT-TMB data timing. 

· Rough timing for the TMB matching window of ALCT vs. CLCT 

· Use TMB logic analyzer and histogramming to debug arrival of signals. 

· Check the DMB readout: do we always have ALCT, CLCT, and CFEB readout? 

· Do we have the same BX numbers for ALCT, CLCT, and CFEB readout? If not, then adjust BX offsets. 

Trigger debugging with two chambers: 

· Make sure we get the same BX numbers for both chambers. 

· What is the fraction of events for which we have really full information (both ALCT, CLCT, and CFEB readouts)? It should be nearly 100% for a good scintillator trigger. 

· Does the MPC receive the muon at the same time from both chambers? Do they have the same BX tag? 

Procedures: 

· Look with a scope at CCB input from scintillators with respect to the 40 MHz clock. Is it centered? If not, add a cable delay to the scintillator trigger so that it is. Do the muons vary with respect to the clock (they shouldn't vary more than 1-2ns)? What is the "jitter" of muon arrival time? Can some actual data be taken on this in order to accumulate a histogram? (GPIB scope readout?) 

· Event displays should look "perfect" when the system is really ready. Both chambers should show the same muon on the event display. 

· TMB logic analyzer events and histograms should also look "perfect" when the system is ready to go. 

· Data dumps should be thoroughly studied before the system is declared to be working properly. 

· Run delay curves of ALCT delay chips with beam in at least 5 places on the chamber: center, and near each of 4 corners. Set all of the ALCT delays to the proper values. 

· Take data and look at timing: the BX of the ALCT should be plotted versus the BX of the scintillator trigger. One convenient place to look is using the logic analyzer display of the TMB. 

· One way to measure efficiencies: use the scintillator trigger to define "events", simply look at the DDU data which has one record for every Level 1 Accept as the denominator. 

· Another way to measure efficiencies: use one chamber as "trigger", use ALCT*CLCT or any other configuration, and look at how often the particular type of data is seen in the other chamber. This should be fairly symmetric between the two chambers if they are operating pretty much the same. 

============================================================

· (Silja) Patterns versus angles in the (,b) scans: plot versus b at the different  settings:

· CLCT and ALCT efficiency

· CLCT pattern

· CLCT and ALCT pattern quality

· (unassigned) Plot CLCT pattern quality versus various trigger thresholds (solve a mystery of missing triggers).

· (Erin project #1) Plot differences between LCTs and the digi LCTs that are simulated from the raw hits readout:

· ALCT discrepancies

· CLCT discrepancies

· Study any discrepancies found versus conditions we changed (particle rate, high voltage, angles, etc.)

· (Brian) look at logic scope data:

· Look at the timing of trigger signals.

· Look at timing versus drift delay.

· (Erin project #2) Look at high voltage scan data (gain of chamber changes quite a bit). Plot versus high voltage:

· CLCT and ALCT efficiency 

· Do we get more di-strip patterns at lower voltages?

· Do other quantities change?

· (Matt) Look at comparator threshold data. Plot versus threshold:

· CLCT efficiency.

· CLCT pattern number.

· Other quantities?

· Look at runs that changed ALCT delay settings in synchronized and unsynchronized beam data, plot:

· efficiency of ALCT bunch crossing. Maximize the statistics.

· Look at runs with various particle rates, plot versus particle rate:

· CLCT and ALCT efficiency.

· Cathode and Anode raw hits readout efficiency.

· CLCT and ALCT pattern qualities.

· CLCT pattern number.

Another list (from original goals list):

· Measure the efficiency for ALCT and CLCT on straight-through tracks. 

· Study, check and measure the ALCT-CLCT matching efficiency 

· Measure ALCT BX identification efficiency and learn how to adjust BXN clock with respect to ALCT timing to maximize this efficiency. This is crucial 
for us to use the ALCT BXN distribution to match the time-gaps with thetime-structure beam. 

· Verify that trigger primitives sent correctly to the MPC. 

A more comprehensive list is: 

1. Single-chamber test goals: 

· Measure the scintillator trigger purity - how? 

· Measure the efficiency for ALCT on straight-through tracks. 

· Measure the efficiency for CLCT on straight-through tracks. 

· Study the ALCT BX identification efficiency: we wish to get the correct BX on 99% or more of the events. Measure the ALCT delay settings needed on different ends of the chamber. Study the time distribution of the ALCT BX versus the scintillator trigger BX. 

· Understand what the width of the CLCT time window needs to be in order to match ALCT, is it +-1 BX or does it need to be wider? 

· Measure the overall chamber trigger efficiency, i.e. requiring ALCT*CLCT*(correct BX). 

· Vary the pretrigger threshold (currently 2 for ALCT, 4 for CLCT) and measure the resulting ALCT-scintillator and ALCT-CLCT relative timing, i.e. make histograms of delta(BX). The ALCT delay chip settings will have to be changed slightly if the ALCT pretrigger threshold is changed. 

· Measure the efficiency for CLCT versus phi angle. Study how the patterns vary with angle. (Is this well described by the Monte Carlo simulation?) 

· Measure the efficiency for ALCT versus theta angle to map out the envelope for high efficiency. (Is this well described by the Monte Carlo simulation?) 

· The efficiency for CLCT should not vary much with theta angle, verify that this is true. Likewise, the efficiency for ALCT should not vary much with phi angle, verify this. 

· Measure efficiencies versus chamber HV. 

· Measure timing versus chamber HV. 

· Measure the rate dependence of efficiencies (and timing?). 

2. Two-chamber test goals: 

· Check the relative ALCT timing between the two chambers. Is it consistent with the rate of BX misidentification for the chambers individually? 

· Measure the 2-chamber "global" ALCT*CLCT trigger efficiency. 

· Check the BX numbers match in the readout at all stages (ALCT, TMB, DMB, DDU). 

· Measure position and angular resolution. 

Make sure that the readout of 2 chambers do not interfere with each other (all data present and valid). ============================================================

